FLINDERS MEDICAL CENTRE DXA STATISTICAL QUALITY CONTROL SPREADSHEET   (DXASQC)

Version 2.8 (Man) , March 2008

John Cormack and Jane Shearer

Medical Imaging Division

Flinders Medical Centre

Adelaide

South Australia 5042

Telephone 618- 8204-4642

E-mail john.cormack@fmc.sa.gov.au

DXASQC is a spreadsheet template designed to plot and assess the statistical significance of drifts in the calibration of  DXA scanners. It uses standard statistical quality control (SQC) techniques to flag QC phantom measurements which indicate that the scanner calibration has drifted, or may be about to drift, outside accepted calibration limits.

The DXASQC program incorporates the following features:

* 	Shewart chart with superimposed moving average and control limits for both Shewhart and moving average.

* 	EWMA chart with control limits .

* 	Precision chart with control limits.

* 	Calculation and display of statistical parameters.

* 	Calculation of drift-corrected values.

* 	Control limits are user selectable to any number of standard deviations - around 2.6, however, seems to be about the optimum.

* 	Fast start-up for the averaged parameters.

* 	Sample size for averaging and precision calculations can be selected by the user - control limits are automatically modified as appropriate.

* 	Initialisation of SQC sequence after calibration.

* 	Visual and audible flagging of "Out of Control" (OOC) conditions. Audible flaggging can be switched off if desired.

* 	Chronological sorting of input data.

* 	Tracking of measured values over any time interval (down to one minute).

* 	Selection of time period for display/printing and data filtering for any combination of OOC conditions.

* 	Selection of a time window on charts plus scrolling forward and backwards in time.

* 	Easy-to-use user interface for data input, editing and deletion.

* 	Values can be read off charts by clicking and pointing with the mouse.

___________________________________________________________________

HARDWARE AND SOFTWARE REQUIREMENTS, INSTALLATION

To run this spreadsheet, you will need EXCEL 97 up for PC. Any PC hardware capable of running EXCEL will do the job. The software may be provided as a zipped file or  in uncompressed format:

Zipped Format

Copy or download  the  zipped file DXASQCv2_8(Man) to your desktop and double click on it – this will unzip the spreadsheet template DXASQCv2_8(Man) and any associated documentation into a folder named DXASQCv2.8(Man).  This folder may then be moved to any desired directory on your machine.

Uncompressed Format

Simply copy or download the folder named DXASQCv2.8(Man) on to  your desktop or any other desired location on your machine.

NOTE: Ensure that macro security in EXCEL (Tools/Options/Security/Macro Security) is set to the "medium" level - you will still receive a security warning on opening the spreadsheet, but you can safely choose to "enable macros"

A Mac version of the package is also available on request.

______________________________________________________

DISCLAIMER NOTICE

Whilst the utmost care has been taken to eliminate "bugs" from this software, the authors cannot guarantee that it is "bug-free". Consequently, they accept no responsibility for loss of data or any other difficulties which may arise from the use of this template. Also, the authors accept no responsibility for the interpretation of the results presented and any clinical actions taken on the basis of these results.

It is strongly advised that a physician or medical physicist with experience in bone densitometry be involved in the interpretation of results obtained using this spreadsheet.

______________________________________________________

Setting up and using the  Spreadsheet package

- NOTE: This spreadsheet is designed for manual entry of phantom data from any densitometry machine. Another  version of the program (DXASQCV2.8(AUTO)) which allows for automatic access of data directly from the bone densitometry database is  available for GE/Lunar machines which use the GE “Encore” bone densitometry software. It will not run, however, on any other type of densitometer, or on GE/Lunar densitometers which do not use “Encore”.

- This spreadsheet is distributed as a template. When you open the original spreadsheet it will generate a copy of itself that you can set up and save for subsequent use. You can generate as many clones as you like from the original, which will remain pristine! You can have as many instances (clones) of the spreadsheet running on as many machines as you like!

- Open the spreadsheet template on any PC that has MS Office installed (this doesn't have to be the PC that runs the software and drives the scanner, but it may be more convenient if it is). Ensure that macro security in EXCEL (Tools/Options/Security/Macro Security) is set to the "medium" level - you will still receive a security warning on opening the spreadsheet, but you can safely choose to "enable macros"

- You may have to select "View: Full Screen" from the main menu to see everything.

- The template will contain some sample data with a simulated 1 SD drift and a simulated change in precision - you can use this data to familiarize yourself with the various functions of the package.

- Enter your own basic input data (hospital/scanner name, serial numbers etc) in the highlighted cells which come up on first running the template.

- To enter your own measured data in the data entry region, simply select and delete the sample data in the yellow data entry region and enter your own - you can do this manually or use the cut and “Copy/Paste Data” button if you have an electronic version of your data available.

- You will need to enter the reference value, precision and drift tolerance. You can enter these directly if you're confident about what they are, but it's generally better to use the automatic set-up routine as follows:

- Click on the  "Set Reference Value" button. The  "Set Precision" and "Set Drift" buttons will be in red text , indicating that you will have to select them next. Accept the default date period (click the cross on the window that appears) and accept the suggested reference value. Repeat this with the "Set Precision" and "Set Drift" buttons. You should end up with three blue buttons again and an approximate set-up. You then have to fine-tune  this process with a more careful selection of the date ranges that you use for these three parameters as follows by again selecting the “Set” buttons (in the following order):

- Set Reference Value: Have a look through the charts and choose a date range that's appropriate for setting the reference value; this range should incorporate at least 25 points and should be in a time period where there is a high confidence that the machine is "in control" (e.g close to a machine calibration), precision is "in control" and there is no appreciable drift. Click the "Set reference value" button, set the dates, and generally follow the prompts. Unless you are absolutely sure about the current reference value, replace it when prompted with the calculated one. The button text will change to blue.

- Set Precision:  Have a look through the charts and choose a date range that's appropriate for setting the precision; this range should incorporate a longish time period ( at least a year if possible) where there precision is "in control" (look at the "Precision" control chart) . Drift doesn't matter too much here. Click the "Set Precision" button, set the dates, and generally follow the prompts. Unless you are absolutely sure about the current precision value, replace it when prompted with the calculated one. The button text will change to blue.

- Set Drift: Have a look through the charts and choose a date range that's appropriate for setting the drift tolerance; this range should incorporate a longish time period ( at least a year if possible) where there precision is "in control" (look at the "Precision" control chart) . In general, it may be a good idea to use the same date range that you used for setting the precision . Click the "Set Drift" button, set the dates, and generally follow the prompts. Unless you are absolutely sure about the current drift tolerance, replace it when prompted with the calculated one. The button text will change to blue.

- If you've done everything correctly and in the right order you should end up with 3 blue “Set” buttons. Generally precision will be around 0.3 to 0.4 %, and drift around 0.2% to 0.3%. If the calculated drift tolerance setting exceeds about 0.4%, it may indicate a drift that will start to impact on measured patient BMD values; or it could mean that "out of control" data points have been used in setting up the reference, precision and drift tolerance values.

- Use the “Save As” menu command  to save a copy of the spreadsheet you've set up with an appropriate name (e.g.  DXASQCMyHospital#1) for subsequent use.

- On subsequent use, simply double click on the spreadsheet you have created to open it, do what you need to do, and  use the “Save” menu command to save the data you have just updated. Insertion of new measured phantom  data can be done easily by clicking on the "Insert Data" button - this takes you to the appropriate place in the database where you can enter the date and time, the measured value and any comments.  If the text in the "Sort Data" button changes to red, click in this button to initiate a chronological sort. The vertical scroll bar at the side of the window allows you to scroll through the entered data. There is room for 2000 entries on the spreadsheet – for additional entries you will have to set up and continue on another sheet (the last 20 or 30 entries from the previous spreadsheet  can be cut and pasted using the “Copy/Paste Data” button to provide some overlap)

______________________________________________________

Output Charts and Data

Four charts are generated (click on the tabs at the bottom of the spreadsheet to display them):

Shewhart

This is a standard Shewhart plot of the raw data and a moving average. The blue points are the raw data and the green line is the moving average. The green line should lie within the inner "tram lines" (control limits) and the blue points within the outer "tram lines". The time window can be selected (enter  the start date and end date you want to use  and click on the "Set Time Window" button. Ensure that the time window will display the data you want to see. The window can be scrolled and the y-axis scaled as desired.

EWMA

The EWMA (Exponentially Weighted Moving Average) chart may be more sensitive to small drifts in calibration, and is included for this reason. Small drifts are more quickly picked up by EWMA than straight Shewhart. For the EWMA, a lambda factor of around 0.15 seems appropriate - higher values will allow faster tracking  of large changes , lower values will allow tracking of smaller changes, but the response will be slower.

Precision

This chart picks up any statistically significant changes in the precision of the measurements (due to changes in machine function, operator changes or other factors).

Moving Average (MA)

This is similar to EWMA, but does not respond as quickly to small drifts. It is, however, an accurate indicator of the overall drift and can, if necessary and appropriate, be used for drift correction.

Measurements which fall outside the control limits on any of these charts will be flagged as "out of control" (OOC) on the main data entry sheet - an audible signal may also be given when "out of control" data is entered.

______________________________________________________

Additional Features and Information

Other functions included in the main sheet include data filtering, extraction of basic statistics for selected data,  setting of statistical control parameters and drift correction. Users are advised to adhere to the default control parameters unless they are experienced with the use of statistical control charts.  A little experimentation using the sample data provided in the template should give users a feel of how things work

A manual containing additional  information and more detailed operating  instructions is in the process of being prepared (but don’t hold your breath!). In the interim, if you have any problems with the installation or use of this program, contact :

John Cormack

Medical Imaging Division

Flinders Medical Centre

E-mail:  john.cormack@fmc.sa.gov.au

Telephone: 618 8204 4642

Fax: 618 8204 5450